Ventriculostomy updates

5 hourly Assessments  :   Remember :  F L I N D 

These are the first thing you should look for when you walk into a room and                  the last thing you notice before you leave !
F luctuation - 

       movement of fluid in the tubing :   fluctuates with each heart beat  : 

       No fluctuation? : Check integrity then: Page Neurosurgery immediately

L evel
     Laser level at “0” , Balance the Bubble , level to Foramen of  Monroe,                             

     Level the black line indicator  to MD order (cm H20)
I ntegrity     

        Stopcocks in correct position           Tubing without kinks 

        No clots                                            Check the site                                       
N euro status

D rainage Color, clarity , amount,

There should be a sign placed above the head of the patient instructing EVERYONE that the level / height /HOB /position of bed SHOULD NOT 

 be changed unless the nurse is present to clamp the ventric                 .

To Level Ventric


Slide laser level into place at the 0 level


          ( Laser’s are NOT Disposable, please return to your units storage area / pyxis when not in use )





BALANCE THE BUBBLE


       MUST DO! !  


	Use the GREY knob to tilt the ventric until the  BUBBLE is  between the two black lines 








3.    Level ventric at  Foramen of Monroe


      (1/2 way between the outer canthus of the eye                     and the ear canal)  


       To activate laser twist silver laser


     (Don’t forget to turn it off when you are done)





Level the cylinder to the MD order  


        *****USE THE BLACK INDICATOR LINE*****


           using the GREEEN scale    cm H20  


       Loosen the white knob to raise or lower the black pointer to the MD order                                  


      (DO NOT USE THE BLUE mmHg SCALE ) 

















NEVER: level the tubing per manufacturers recommendation


Codman EDS 2 Manual


  " Note The Drip Chamber pressure reference line corresponds to the fluid outlet at the antireflux valve within the Drip Chamber. The inlet tubing above the antireflux valave has no effect on the pressure setting because both sides of the tubing are equidistant to its peak from the pressure setting arrow."











